STEP - ASIC IDefX BD
We discovered an issue with some channels(pixels) in the two ASIC IDefX BD‘s used in the
SupraThermalElectronProton Instrument onboard SOLarOrbibter.
Pulse Hight Analysis (PHA) data show „ghost“ peaks on several pixels that are linked to
precursor events that exceed the nominal energy range of STEP.
It seems that for very energetic events some of the chanels are not at zero after event
processing, and instantly trigger again on the declining flank of the event.
To illustrate the problem :
- Expected form of observed spectra (Pages 2,3,4)
- Comparison to conspicuous Pixel (Pages 5,6,7)
- Normalised PHA spectra of all pixels for both heads (Page 8,9)
- Delta times to last processed events (Pages 10,11)
- Spectra for all pixel for delta times < 20 µs and > 20 µs (Pages 12,13)
- Energy spectra of precoursor events (Page 14)
- Spektrum for pixel 1 of head 1 decomposed (Page 15)

---
STEP - SETUP

---
STEP - SETUP

---
STEP - SETUP

---
Head 0, Pixel 1 ADC Spectrum
INT Pixel 1 ENERGY

---
Head 0 Pixel 1 Energy Spectrum
[keV]

---
Head 0 Pixel 1 Energy Spectrum Copper Line
[keV]

---
Heads 0,1 Pixel 1 ADC Spectra

---
Heads 0,1 Pixel 1 Energy Spectra
[keV]

---
Heads 0,1 Pixel 1 Energy Spectra Copper Line
„Ghost“ Peak
[keV]

---
Head 0 Pixel all Energy Spectra Normed Linear
„Ghost“ Peaks
[keV]

---
Head 1 Pixel all Energy Spectra Normed Linear
„Ghost“ Peaks
[keV]

---
MAG Pixel ALL ENERGY DT > 20 µs
Head 1 Pixel all DeltaTimes
Max. DeltaTime
(Overflow bin)
Unexpected
[µs]

---
MAG Pixel ALL ENERGY DT > 20 µs
Head 1 Pixel all short DeltaTimes
Pixels with „Ghost“ Peaks
All Pixels are similar
Pixels without
 „Ghost“ Peaks
[µs]

---
MAG Pixel ALL ENERGY DT > 20 µs
Head 1 Pixel all short DeltaTimes
Fractions of dtimes < 20 µs
Pixels with „Ghost“ Peaks
All Pixels are similar
Pixels without
 „Ghost“ Peaks
[µs]

---
MAG Pixel ALL ENERGY DT > 20 µs
Head 1 Pixel all short DeltaTimes
Fractions of dtimes < 20 µs
Timeseries „healthy“ pixel
Pixels with „Ghost“ Peaks
All Pixels are similar
Pixels without
 „Ghost“ Peaks
[µs]

---
MAG Pixel ALL ENERGY DT > 20 µs
Head 1 Pixel all short DeltaTimes
Fractions of dtimes < 20 µs
Timeseries „strange“ pixel
Timeseries „healthy“ pixel
Pixels with „Ghost“ Peaks
All Pixels are similar
Pixels without
 „Ghost“ Peaks
[µs]

---
Head 1 Pixel all Energy Spectra DeltaTimes < 20 µs
„Ghost“ signals look like
 squeezed spectra
Expected shape for high-flux
situations (powerlaws)
[keV]

---
Head 1 Pixel all Energy Spectra DeltaTimes > 20 µs
No „Ghost“ Peaks !
[keV]

---
Head 1 Pixel all Energy Spectrum Precursor Events DeltaTimes < 20 µs
MAG Pixel ALL ENERGY DT > 20 µs
Normal Pixels
Strange Pixels
[keV]

---
Head 0 Pixel 1 Energy Spetra for : DeltaTtime < 20 µs, > 20 µs, all
MAG Pixel ALL ENERGY PREVIOUS EVENT DT < µs
MAG Pixel ALL ENERGY DT > 20 µs
> 20 µs
all (sum)
< 20 µs
[keV]