## PF RADBC for ref.
MSL RAD FEE Board-Test
RADBC v03 for refurbished BC-stack, no detetctors.
DUT Configuration: | [8 empty] | Date: | 2007-06-20
Mark items that are part of the DUT with a 1. The expected supply current will be calculated.
[7 empty] | Icc/Iss | Ifet
| Boards |  |  | CSA | FSH |  | 0.4 | 4 |  | Detectors
| A |  |  | 2 | 4 |  | 0 | 0
| BC | 1 |  | 6 | 6 |  | 4.8 | 24 |  |  | H | M | L
| DH |  |  | 1 |  |  | 0 | 0 |  | A
| DM |  |  | 1 |  |  | 0 | 0 |  | B
| DL |  |  | 1 |  |  | 0 | 0 |  | C
| E |  |  | 3 | 3 |  | 0 | 0 |  | D
| F |  |  | 3 | 4 |  | 0 | 0 |  | E
| Tflex | [8 empty] | F
|  |  |  | 6 | 6 |  | 4.8 | 24
Ohmmeter Test | [8 empty] | Date: | 2007-06-20
Measure with a (yellow Fluke) Ohmemeter between the indicated pin and ground, both polarities.
| DVM type: |  |  | Fluke 177 |  |  |  | expect:
| + on Vbias |  |  |  | inf | MOhm |  | 5
| -- on Vbias |  |  |  | inf | MOhm |  | inf
| + on Vfet |  |  |  | 133.9 | Ohm |  | 133.3333
| -- on Vfet |  |  |  | 134 | Ohm |  | 133.3333
| + on Vcc |  |  |  | 2.05 | MOhm |  |  | 6 MOhm range
| -- on Vcc |  |  |  | 65.19 | kOhm
| + on Vss |  |  |  | 65.72 | kOhm
| -- on Vss |  |  |  | 2.084 | MOhm |  |  | 6 MOhm range
Bias Insulation Test | [8 empty] | Date: | 2007-06-20
Apply -70V to Vbias, slowly ramping up, measure the current.
Ramp the voltage slowly, to protect the jFETs.
| Ibias |  | @ |  | -70 | V | = | < 4 | nA
Supply Current Test | [8 empty] | Date: | 2007-06-20
Apply power to the board, and measure supply currents
expected:
| Icc |  | @ |  | 5 | V | = | 6.04 | mA | 4.8 |  | Hot lab (sunshine)
| Iss |  | @ |  | -5 | V | = | -6.07 | mA | 4.8 |  | Hot lab (sunshine)
| Ifet |  | @ |  | 5 | V | = | 23.84 | mA | 24
DC Output levels | [8 empty] | Date: | 2007-06-22
Measure Output Voltages with a DVM
expect:
A | A1H |  |  |  |  | mV |  | 250 | 50
| A1L |  |  |  |  | mV |  | 50 | 20
| A2H |  |  |  |  | mV |  | 250 | 70
| A2L |  |  |  |  | mV |  | 50 | 20
BC | BH |  |  |  | 152 | mV |  | 250 | 70
| BL |  |  |  | 42 | mV |  | 50 | 20
| CH |  |  |  | 163 | mV |  | 150 | 30
| CL |  |  |  | 43 | mV |  | 50 | 20
| C2 |  |  |  | 180 | mV |  | 150 | 30
| F1 |  |  |  | 900 | mV |  | 800 | 150
DH | DH |  |  |  |  | mV |  | -250 | 50
DM | DM |  |  |  |  | mV |  | -250 | 50
DL | DL |  |  |  |  | mV |  | -250 | 50
E | EH |  |  |  |  | mV |  | 250 | 70
| EM |  |  |  |  | mV |  | 120 | 40
| EL |  |  |  |  | mV |  | 50 | 20
F | F2 |  |  |  |  | mV |  | 800 | 150
| DH |  |  |  |  | mV |  | 250 | 70
| DM |  |  |  |  | mV |  | 150 | 30
| DL |  |  |  |  | mV |  | 50 | 20
Bias Insulation Test | [8 empty] | Date: | 2007-06-22
Apply -70V to Vbias, slowly ramping up, measure the current.
Ramp the voltage slowly, to protect the jFETs.
| Ibias |  | @ |  | -70 | V | = | < 4 | nA
Supply Current Test | [8 empty] | Date: | 2007-06-22
Apply power to the board, and measure supply currents
expected:
| Icc |  | @ |  | 5 | V | = | 4.8 | mA | 4.8
| Iss |  | @ |  | -5 | V | = | -4.48 | mA | 4.8
| Ifet |  | @ |  | 5 | V | = | 23.91 | mA | 24
NCR | [8 empty] | Date: | 2007-06-21
Symptoms:
Stimulation only two of three F1 inputs produces no output signal.
DC output level of about 250mV, expect about 800mV.
Diagnosis:
Two wires SFb_R2 and SFc_R2 were misplaced.
Remedy:
replace wire correctly.
Fixed:
Done by SiB 2007-06-21
NCR | [8 empty] | Date: | 2007-06-22
Symptoms:
B-channel output amplitude to high by about factor 2.5.
Diagnosis:
(Guess) Preamp feedback capacitor LB_C2 = 3.3pF, instead of 8.2 pF.
Remedy:
Replace LB_C2 by 8.2pF capacitor.
Fixed:
Pulser Stimulation test | [8 empty] | Date: | 2007-06-22
Solder pulse injector wires to the preamp inputs, and observe output pulse shapes and amplitudes.
F1 output pulse with stimulation on F1a with 100mV*1pF, measured with DIRENA.
B, C, and C2 output pulse measured with DIRENA.
B, C, and C2 output pulse height spectra measured with DIRENA.
Observe that the BL and CL amplitudes are similar, suggesting that the B feedback capacitor is wrong.
F1 output pulse height spectra, with stimulation all the three CSA, measured with DIRENA.
Images | [8 empty] | Date:
http://www.ieap.uni-kiel.de/et/msl/pictures/photos/rsh/Tests/PFBCR/index.html

## Template
MSL RAD FEE Board-Test
? what is being tested
DUT Configuration: | [7 empty] | Date:
Mark items that are part of the DUT with a 1. The expected supply current will be calculated.
[6 empty] | Icc/Iss | Ifet
| Boards |  | CSA | FSH |  | 0.4 | 4 |  | Detectors
| A |  | 2 | 4 |  | 0 | 0
| BC |  | 6 | 6 |  | 0 | 0 |  |  | H | M | L
| DH |  | 1 |  |  | 0 | 0 |  | A
| DM |  | 1 |  |  | 0 | 0 |  | B
| DL |  | 1 |  |  | 0 | 0 |  | C
| E |  | 3 | 3 |  | 0 | 0 |  | D
| F |  | 3 | 4 |  | 0 | 0 |  | E
| Tflex | [7 empty] | F
|  |  | 0 | 0 |  | 0 | 0
Ohmmeter Test | [7 empty] | Date: | 0
Measure with a (yellow Fluke) Ohmemeter between the indicated pin and ground, both polarities.
| DVM type: |  | Fluke 177 |  |  |  | expect:
| + on Vbias |  |  |  | MOhm |  | #DIV/0!
| -- on Vbias |  |  |  | MOhm |  | inf
| + on Vfet |  |  |  | Ohm |  | #DIV/0!
| -- on Vfet |  |  |  | Ohm |  | #DIV/0!
| + on Vcc |  |  |  | MOhm |  |  | 6 MOhm range
| -- on Vcc |  |  |  | kOhm
| + on Vss |  |  |  | kOhm
| -- on Vss |  |  |  | MOhm |  |  | 6 MOhm range
Bias Insulation Test | [7 empty] | Date: | 0
Apply -70V to Vbias, slowly ramping up, measure the current.
Ramp the voltage slowly, to protect the jFETs.
| Ibias | @ |  | -70 | V | = |  | nA
Supply Current Test | [7 empty] | Date: | 0
Apply power to the board, and measure supply currents
expected:
| Icc | @ |  | 5 | V | = |  | mA | 0
| Iss | @ |  | -5 | V | = |  | mA | 0
| Ifet | @ |  | 5 | V | = |  | mA | 0
DC Output levels | [7 empty] | Date: | 0
Measure Output Voltages with a DVM
expect:
A | A1H |  |  |  | mV |  | 250 | 50
| A1L |  |  |  | mV |  | 50 | 20
| A2H |  |  |  | mV |  | 250 | 70
| A2L |  |  |  | mV |  | 50 | 20
BC | BH |  |  |  | mV |  | 250 | 70
| BL |  |  |  | mV |  | 50 | 20
| CH |  |  |  | mV |  | 150 | 30
| CL |  |  |  | mV |  | 50 | 20
| C2 |  |  |  | mV |  | 150 | 30
| F1 |  |  |  | mV |  | 800 | 150
DH | DH |  |  |  | mV |  | -250 | 50
DM | DM |  |  |  | mV |  | -250 | 50
DL | DL |  |  |  | mV |  | -250 | 50
E | EH |  |  |  | mV |  | 250 | 70
| EM |  |  |  | mV |  | 120 | 40
| EL |  |  |  | mV |  | 50 | 20
F | F2 |  |  |  | mV |  | 800 | 150
| DH |  |  |  | mV |  | 250 | 70
| DM |  |  |  | mV |  | 150 | 30
| DL |  |  |  | mV |  | 50 | 20